HPV Technology validation study for processing self-collected swab samples

August 2022

Background

All New Zealand laboratories processing clinician-taken LBC vial samples for HPV testing are required to use HPV technologies that have achieved international validation (FDA approval or CE marking) for processing these samples for HPV testing. With self-collected swab samples, the sampling device and the sampling method differs from those used in the clinical trials on which international validation for HPV testing on LBC vial samples is based. Validating swab-collected samples for HPV testing internationally requires new clinical validation trials and international endorsement by either the United States Food and Drug Administration (FDA) or the Conformité Européenne.

Some HPV test technologies (e.g. BD Onclarity) have already obtained international endorsement for processing self-collected vaginal swab samples for HPV testing. No further validation is required to  process swab samples for HPV testing in New Zealand using these test technologies, as long as the same sampling device is used as was used in the international trials on which the endorsement was based.

Other companies have not achieved international endorsement for processing self-collected swab samples for HPV testing with their HPV technologies so currently could not be used in New Zealand when the NCSP introduces self-sampling within the HPV primary screening programme in mid-2023. This is also an issue now for some laboratories in relation to self-testing research studies currently underway in New Zealand.

The NCSP have developed this protocol to support laboratories to validate their HPV test technology to process swab-collected samples based on validation protocols devised and used in Australia for this purpose. The current study is a collaboration between the NCSP and Southern Community Laboratories (SCL, Dunedin) with support from an external biostatistician and supervision from a sector-based governance group.

The Protocol

The validation process will be run as a collaboration between the NCSP and the laboratory (SCL Dunedin). The two Clinical Leads in the NCSP (Pathology and Colposcopy) will arrange the sample collection and the final results analysis and the Lead Molecular Scientist at the laboratory involved will oversee the HPV testing and data collection. The NCSP will fund the additional HPV test costs. A governance group of New Zealand Lead molecular scientists, a microbiologist (Dr Anja Werno) and an international expert molecular scientist (Associate Professor David Hawkes, VCS Pathology (a division of the Centre for the Prevention of Cervical Cancer), Melbourne) will oversee the project. Wendy Dallas-Katoa (Māori Advisor) has reviewed the study documentation and will join the governance group from August 2022. She is available for further consultation should this be needed during the course of the study.

Main centre colposcopy units will collect 320 paired self-collected swab samples and clinician-taken LBC samples for HPV testing.

A sample size of 313 paired samples will be sufficient to estimate HPV test sensitivity using self-collected swabs versus LBC samples at an acceptable level of precision (based on an HPV positivity rate of at least 40% in samples collected at colposcopy). To cover the possibility that a few samples may not be suitable for processing, 320 paired samples will be collected.

- 320 paired samples will be collected for each laboratory undergoing a validation trial. Colposcopists have agreed to collect the samples free of charge.

- Ethnicity data will be collected for study participants from the laboratory request form and included in the report analysis.

- The clinician-taken samples will be part of the patient record. As cytology will be requested on many LBC samples, this needs to be done in a way that does not compromise the cytology result.  For this reason, LBC specimens will be processed for cytology first then the remaining specimen used for HPV Testing. Some specimens may have exhausted the specimen during the cytology processing so these cases would not be included in the final number of samples processed for the study.

- The clinician-taken samples will be processed as “specialist testing”, as already approved in the current NCSP.

  - Dr Howard Clentworth (NCSP Clinical Lead, Colposcopy) will work with lead colposcopists in the participating DHB hospital colposcopy clinics.

  - Written consent from women will be obtained.

  - A Participant information Sheet, a consent form, self-collection instructions and a laboratory request form identifying HPV validation samples, will be provided to colposcopy units

- Ethics approval will be obtained from the Health and Disability Ethics Committee  (HDEC) before commencing the study

- International Accreditation New Zealand (IANZ) confirm that they don’t need to provide approval as they are an accreditation body, not a technology-approval body.

- The laboratory will be responsible for performing the tests and reporting on the results to the NCSP.

- External data/statistical experts will assist and oversee data collation and analysis

Note:

- Once a laboratory has validated one type of HPV test technology, it will be possible to run a comparative in-house validation exercise for a different type of HPV test technology for processing self-collected swab samples. This validation exercise requires (as long as the same device/media combination is used) processing of 30 positive self-collected samples on the validated platform and on the new platform requiring validation. The Australian National Pathology Accreditation Advisory Council (NPAAC) protocol(2) provides the basis for this test number.

Sample collection. Patients attending colposcopy clinics will be asked to participate and sign the written consent form. They will be given verbal and written instructions for self-collection of a swab sample. They will then continue with their  colposcopy during which the colposcopist will take an LBC sample for HPV testing, and cytology as required for clinical management. The request form will identify the samples as part of the HPV self-testing validation.

Laboratory processing, reporting and sample disposal.

Each self-collected swab sample will be taken using a flocked swab, such as the Copan FLOQSwab 552C.80 and either sent to the laboratory as a dry swab (preferable) or placed into a fluid medium depending on the requirements of the HPV testing platform. The system used must replicate the collection system that will be used routinely when HPV self-collected swab samples are used for HPV reporting. Samples will be tested on the HPV platform  requiring validation at the laboratory. Results will be reported as HPV Not Detected, HPV Detected: HPV 16, HPV 18, and/or HPV Other or HPV Invalid, as in regular practice.

- All samples will be stored securely in the laboratory for at least one month and until the study is completed. A karakia will be available at the time of disposal, to accommodate the wishes of wāhine of Māori ethnicity.

- Participants will be able to request that their samples or data are withdrawn from the study, should they decide to do so after participating. Notification to prevent processing of samples will need to occur within 24 hours of the sample being taken, as processing would occur within this time frame.

- Participants will be able to request that the results of both tests will be disclosed to them, should they wish this to occur. All clinician-taken sample results will be known to the treating colposcopist and the vaginal swab sample results will be made available through the treating colposcopist, on request.

- Participants will be able to request a copy of the final study report by making their request known through their treating colposcopist.

Managing the results generated: The NCSP has confirmed that the following will apply:

- self-collected swab HPV results will be held in the reporting laboratory and not reported to the NCSP Register, as the testing will not be validated at the time it was conducted.

- clinician-taken HPV results will be reported in the usual way as specialist testing (i.e. within current guidelines) with a copy to primary care,  and will contribute to the participant’s NCSP record and to clinical management. All samples will be taken by colposcopists who will manage the results so there will be no management implications for the NCSP Register.

Governance. A clinical governance group has been established to confirm approval of the validation protocol. This group consists of

- five lead molecular scientists for laboratories performing HPV testing,

- Dr Anya Werno from Canterbury Health Laboratories,

- Associate Professor David Hawkes, VCS Pathology (a division of the Centre for the Prevention of Cervical Cancer), Melbourne (international expert/reviewer),

- Drs Howard Clentworth and Margaret Sage (NCSP Clinical leads)

- Nevena Novakovic (NCSP Senior Portfolio Manager).

- Wendy Dallas- Katoa ( Māori advisor) ( from August 2022)

The governance group will approve the validation methodology.  Assoc Professor Hawkes will provide expert advice and will review the laboratory findings. The results  will remain confidential to the NCSP (and the international reviewer) for reasons of commercial sensitivity but will also be released to the manufacturer of the HPC testing platform, as evidence of validation for using their technology  to process self-collected swab samples for HPV testing.

NCSP Clinical Governance Group(NCSP COG) has approved the validation protocol subject to further expert peer review within New Zealand.

Analysis of results

HPV positivity for HPV-16, HPV-18 and HPV-Other or any oncogenic HPV type will be recorded for each pair of samples and the results of self-collected swab samples compared with the clinician-taken sample HPV test results following the approach used by Saville et al. In brief, the number and proportion of samples that have invalid or failed tests results will be calculated by testing method and differences assessed using Bowker’s test of symmetry for paired samples. The accuracy of test results from valid self-collected samples will be compared with test results from paired valid clinician-taken samples by calculating the sensitivity, specificity, and observed agreement with 95% Wilson score binomial confidence intervals for detecting HPV16, HPV18, other oncogenic HPV (non-HPV 16/18), and any oncogenic HPV types (HPV31/33/35/39/45/51/52/56/58/59/66/68).  We will also estimate the sensitivity of self- and clinician-collected samples using an alternative reference standard where any positive result from either sampling method is indicative of infection.

|  |  | Clinician-taken sample (reference) |  |
| --- | --- | --- | --- |
|  |  | HPV + | HPV - |
| Self-taken sample | HPV + | A | B |
|  | HPV - | C | D |

Sensitivity = A / (A + C)

Specificity = D / (B + D)

Observed agreement = (A + D) / (A + B + C + D)

The overall level of agreement will also be assessed by calculating Gwet’s AC1 coefficient for interrater agreement, where AC1 ≤ 0.20 is considered as poor, 0.21 ≤ AC1 ≤ 0.40 as fair, 0.41 ≤ AC1 ≤ 0.60 as moderate, 0.61 ≤ AC1 ≤ 0.80 as good, and AC1 ≥ 0.81 as very good agreement.

Assuming an HPV positivity rate of at least 40% in samples collected at colposcopy, a sample size of 313 will be of sufficient size to estimate test sensitivity with 95% confidence interval of no wider than plus or minus 8%. This sample size will also allow the estimation of Gwet’s AC1 coefficient with confidence intervals no wider than plus or minus 8%.

Progress to the governance group will be advised as the project proceeds.

Final Study report

The final report will be complied by the Lead Molecular Scientist at the participating laboratory in conjunction with the biostatistician and reviewed by the Study Governance Group and the NCSP before sign-off. The final report will be held by SCL and made available to the Study Governance group, the NCSP and the commercial company responsible for the HPV test technology used, as these groups require confirmation of validation of the use of self-collected vaginal swab samples for HPV testing with the test technology used at the specific laboratory participating in the study.

References

1. J Clin Virol. 2020 Jun;127:104375. doi: 10.1016/j.jcv.2020.104375. Epub 2020 Apr 18.

- Analytical performance of HPV assays on vaginal self-collected vs practitioner-collected cervical samples: the SCoPE study

M Saville 1, D Hawkes 2, Mht Keung 3, Elo Ip 3, J Silvers 4, F Sultana 5, M J Malloy 5, L S Velentzis 6, K Canfel L 7, C D Wrede 8, Jml Brotherton 9

2. NPAAC protocol attached.

Dr Margaret Sage, Clinical Lead, Pathology National Cervical Screening Programme (NCSP)

Dr Jonathon Williman, Biostatistician and Senior Research Fellow (Otago University)

Dr Jenny Grant, Lead Molecular Scientist, Southern Community Laboratories (SCL)